## Slide 1: Sources of Error in Quantitative MRI (qMRI) Parameter Estimation A Signal Processing Perspective
Joshua D. Trzasko, Ph.D.
Assistant Professor of Radiology and Biomedical Engineering
Associate Consultant - Department of Radiology (Medical Physics)
Mayo Clinic, Rochester, MN, USA
ISMRM Virtual Meeting - Quantitative MR Study Group
September 2018

## Slide 2
Speaker Name: Joshua D. Trzasko
I have no financial interests or relationships to disclose with regard to the subject matter of this presentation.
Declaration of Financial Interests or Relationships

## Slide 3: Outline
What is Quantitative MRI (qMRI)?
qMRI: a Signal Processing Perspective
Sources of Error in qMRI
  Noise Model
  Signal Model
  Fitting Strategy
Summary and Closing Remarks
Notes: Broaden persepctive
Unifying theme

## Slide 4: Outline
What is Quantitative MRI (qMRI)?
qMRI: a Signal Processing Perspective
Sources of Error in qMRI
  Noise Model
  Signal Model
  Fitting Strategy
Summary and Closing Remarks
Notes: Broaden persepctive
Unifying theme

## Slide 5: What is Quantitative MRI (qMRI)?
qMRI is any MRI application where the goal is to describe some property of the imaged object using meaningful physical or biological units.

## Slide 6: Why is qMRI Clinically Important?
S. Venkatesh et al. JMRI 37.3 (2013): 544–555. PMC. Web. 18 Sept. 2018.
qMRI offers several key advantage over routine anatomical imaging:
Simplified radiological interpretation
Increased diagnostic confidence
Improved detection and staging of disease, and monitoring of treatment
Example: Liver fibrosis staging via MR Elastography (MRE)
Easier longitudinal and cross-sectional studies
More efficient clinical workflows
Potential for abbreviated exams

## Slide 7: qMRI Categories
S. Teipel et al., The Lancet. Neurology 14 10 (2015): 1037-53.
J. Schwimmer et al. Hepatology 61.6 (2015): 1887-1895.
Structural
Physiological
There are MANY different quantitative applications utilized both clinically and in research settings. These typically fall into one of two categories:
PDFF
In this lecture, we’ll focus exclusively on physiological qMRI

## Slide 8: Example Physiological qMRI Applications
A non-inclusive list*:
*apologies for unintentional omissions
Functional MRI (fMRI)
Flow
T1/T2 Mapping
Pharmacokinetics
Fat+Water Decomposition
Elastography
Diffusion
Perfusion
Thermometry
Susceptibility Mapping (QSM)
Metabolite Concentration

## Slide 9: How is qMRI Data Acquired?
Due to a variety of physical (e.g., receiver gains) and physiological factors (e.g., patient motion), MR image intensity values are not intrinsically quantitative….
However, we typically have (via physics) a solid understanding of the causal factors of an image’s appearance; e.g., in spoiled gradient echo (SPGR) MRI, a pixels intensity is modeled as:
By repeating a scan with different acquisition parameters (e.g., flip angle), the individual contributions of underlying variables (e.g., T1) are revealed:
5o
10o
15o
20o
25o
30o
Arbitrary Scaling Factor (∝ PD)

## Slide 10: How is qMRI Data Processed?
qMRI data is typically processed by:
5o
10o
15o
20o
25o
30o
Forming a signal model
Identify known vs. unknown variables
Defining a fitting strategy
Evaluating performance

## Slide 11: Outline
What is Quantitative MRI (qMRI)?
qMRI: a Signal Processing Perspective
Sources of Error in qMRI
  Noise Model
  Signal Model
  Fitting Strategy
Summary and Closing Remarks
Notes: Broaden persepctive
Unifying theme

## Slide 12: The Challenge of Being Quantitative
Acquisition settings aside, there will be different ways in which parameters can be estimated from qMRI data sets – how do we know which is the “best”?
?
T1 Map - Method 1
T1 Map - Method 2
To fully answer this question, we must first understand why different methods yield (sometimes significantly) different results.

## Slide 13: Sources of Error in qMRI Data Processing
qMRI processing methods each rely on different assumptions about the data and/or fitting method. If these assumptions are violated, errors will form in the parameter estimates. Common error sources include:
Incomplete/inaccurate noise model
Incomplete/inaccurate signal model
Unstable fitting strategy
Insufficient data diversity*
vs
R1 Map (w/o B1+ Correction)
B1+ Map
R1 Map (w/ B1+ Correction)
*Will not be covered in detail here

## Slide 14: qMRI Processing is a Model Fitting Problem
qMRI parameter estimation – like image reconstruction and artifact correction – is fundamentally a parametric model fitting problem. As such, established statistical signal processing tools can be used to both develop and rigorously characterize/optimize qMRI method performance.
Image Reconstruction
qMRI

## Slide 15: Key Tools for qMRI Parameter Estimationr
Hey Ron, want to work together?
Sounds like a jolly good time, Felix!
Felix Bloch, 1952
Ronald Fisher, 1956
George Dantzig, 2001
You’re probably going to need a little help….
MR Physics
Statistical Signal Processing
Nonlinear Optimization Theory
Notes: Maximize current systems
Value to patient – lower cost
Speed
Better
Diminishing healthcare reimbursements

## Slide 16: Introduction to Estimation Theory
Given a data set, g, parameter vector θ, and random variable model, A{}:
our goal in qMRI is to design a estimator (mathematical functional), H{}, that we will apply to the measured data, g, to obtain an estimate of θ:
We have total freedom in designing H{}; however, a useful estimator will have:
Reminder: since g is random, any estimate derived from it is itself random.
Bias = Accuracy-1
Variance = Precision-1

## Slide 17: Mean Square Error (MSE)
Both the bias and variance of any estimator parameter are important – how can measure both quantities simultaneously?
MSE is the standard estimator performance metric:
Low Accuracy
Low Precision
Low Accuracy
High Precision
High Accuracy
Low Precision
High Accuracy
High Precision

## Slide 18: Outline
What is Quantitative MRI (qMRI)?
qMRI: a Signal Processing Perspective
Sources of Error in qMRI
  Noise Model
  Signal Model
  Fitting Strategy
Summary and Closing Remarks
Notes: Broaden persepctive
Unifying theme

## Slide 19: Noise in MRI
All MRI data – whether raw or reconstructed – contains “noise”, or random fluctuations from thermal interactions in our bodies and scanner electronics.
Noise in raw MRI data is proper complex AWGN*. However, this is generally no longer true once data has been processed; e.g., computing a magnitude:
*Phased array data may be correlated across the channel dimension
Inaccurate noise modeling (e.g., ignoring it) is one of the greatest sources of error in qMRI parameter estimation.
AWGN (Real, Imag)
Rician

## Slide 20: Example: T2 Mapping
T2 mapping is often considered one of the “easiest” qMRI applications – but is this just a case of “ignorance is bliss”? Consider the following decay curve:
m0 = 1;
TE = linspace(5,500,12)'*1e-3; %ms
T2 = 70e-3; % ms
g0 = m0*exp(-TE/T2);

## Slide 21: Example: T2 Mapping
Now let’s add some complex AWGN and compute the signal’s magnitude. Repeat this 1000x at 1000 different noise levels.
m0 = 1;
TE = linspace(5,500,12)'*1e-3; %ms
T2 = 70e-3; % ms
g0 = m0*exp(-TE/T2);
T = 1000; Ns = 1000; Nt = length(TE);
sigma = linspace(1e-4,0.05,Ns)';
X.raw = zeros([Nt Ns T]);
for s=1:Ns
for t=1:T
X.raw(:,s,t) = g0 + sigma(s)*(randn([Nt 1]) + 1i*randn([Nt 1]));
end
end
X.mag = abs(X.raw);
*Only a subset of instances visualized

## Slide 22: Example: T2 Mapping
Compute the mean across trials. Observe that as noise increases, the exponential (erroneously) appears to decay more slowly – this is noise bias.
m0 = 1;
TE = linspace(5,500,12)'*1e-3; %ms
T2 = 70e-3; % ms
g0 = m0*exp(-TE/T2);
T = 1000; Ns = 1000; Nt = length(TE);
sigma = linspace(1e-4,0.05,Ns)';
X.raw = zeros([Nt Ns T]);
for s=1:Ns
for t=1:T
X.raw(:,s,t) = g0 + sigma(s)*(randn([Nt 1]) + 1i*randn([Nt 1]));
end
end
X.mag = abs(X.raw);
Y = mean(Y,3);

## Slide 23: Example: T2 Mapping
Monte Carlo simulation can also be used to reveal the bias-variance behavior of specific estimators; e.g., linearized method-of-moments for T2 estimation:
Mean
Mean+Std. Dev
Mean-Std. Dev

## Slide 24: Example: T2 Mapping
For real-world SNR levels, noise-based errors in standard T2 fitting results are dominated by lost accuracy (bias) rather than lost precision (variance).

## Slide 25: Noise Affects Every Method Differently
Trzasko et al. MRM 69(6):1787-1794, 2013.
200 ms
1000 ms
4000 ms
Volunteer #1
Joint
RSS Composite
Debiased RSS
Channel-by-Channel
Linearized Approx.
NLLS
Below are the results of 8 different T1 mapping variants applied to the same (complex) multi-channel VFA-SPGR data set – note how performance varies.

## Slide 26: Making Informed Tradeoffs
[a] Linearized Approx. – Joint [e] NLLS – Joint
[b] Linearized Approx. – Channel-by-Channel [f] NLLS – Channel-by-Channel
[c] Linearized Approx. – RSS Composite [g] NLLS – RSS Composite
[d] Linearized Approx. – Debiased RSS Composite [h] NLLS – Debiased RSS Composite
Joint nonlinear least squares (NLLS) fitting is expected to yield the lowest MSE results for this application – how do the other methods perform relatively?
Trzasko et al. MRM 69(6):1787-1794, 2013.

## Slide 27: Explaining the Unexpected
Noise is also commonly the source of unanticipated changes in results that can occur following qMRI protocol changes; e.g., DWI with log-linear fitting:
Saritas et al. IEEE TMI 30:424-437, 2011
Trzasko et al. Proc. ISMRM 2018
Notes: Changing b-values, NEX
For the first protocol (P1), TR/TE=3700/82.8ms, b=[100;1000;1600]s/mm2, and NEX=[5;16;16]; and for the second protocol (P2), TR/TE= 4050/69.3ms, b=[100;400;1000]s/mm2, and NEX=[3;5;14].
Complex averaging

## Slide 28: Regularization: Trading Bias for Variance
Estimation performance can often be improved by leveraging auxiliary prior knowledge about the target quantity, such as spatial smoothness or sparsity. In such scenarios, one is intentionally biasing the result in a specific direction.
Prior knowledge is typically incorporated into a fitting procedure by (additively) augmenting its cost function with an regularization/penalty term; e.g.,
If the data fidelity term coincides with that of the Maximum Likelihood estimator, this strategy is often called maximum a posteriori (MAP) estimation.
Data Fidelity
Penalty
Mixing
Parameter

## Slide 29: Example: Diffusion Weighted Imaging
The signal in a diffusion weighted imaging (DWI) can be modeled as , where the apparent diffusion coefficient (ADC) is the primary quantity of interest.
E. Borisch et al., Proc. ISMRM 2018
K. Keenan et al. MRM 79(1):48–61, 2018
Standard (MOM)
Maximum Likelihood

## Slide 30: Example: Diffusion Weighted Imaging
The signal in a diffusion weighted imaging (DWI) can be modeled as , where the apparent diffusion coefficient (ADC) is the primary quantity of interest.
E. Borisch et al., Proc. ISMRM 2018
K. Keenan et al. MRM 79(1):48–61, 2018
Standard (MOM)
Regularized

## Slide 31: Example: MR Elastography
Regularized harmonic estimation improves MRE displacement estimates, which in turn yields more accurate/precise tissue stiffness estimates and pathology localization; e.g., vestibular schwannoma (skull base tumor):
Trzasko et al. Proc. ISMRM 2016

## Slide 32: Minimizing Error Propagation
Advanced image reconstruction methods often rely on quantitative reference information like ∆B0 maps. Any errors in these quantities will propagate into and compromise the quality of the final reconstruction results; e.g., spiral:
Field Map
No ORC
ORC + Standard
ORC + Robust
Gridding Recon
Sutton et al. IEEE TMI 22(2):178-188, 2003
Trzasko et al. Proc. RSNA, SSQ17-04, 2012
Tao et al. Medical Physics 42(12):7190–7201, 2015.

## Slide 33: Minimizing Error Propagation
Advanced image reconstruction methods often rely on quantitative reference information like ∆B0 maps. Any errors in these quantities will propagate into and compromise the quality of the final reconstruction results; e.g., spiral:
Field Map
No ORC
ORC + Standard
ORC + Robust
Gridding Recon
Sutton et al. IEEE TMI 22(2):178-188, 2003
Trzasko et al. Proc. RSNA, SSQ17-04, 2012
Tao et al. Medical Physics 42(12):7190–7201, 2015.

## Slide 34: Outline
What is Quantitative MRI (qMRI)?
qMRI: a Signal Processing Perspective
Sources of Error in qMRI
  Noise Model
  Signal Model
  Fitting Strategy
Summary and Closing Remarks
Notes: Broaden persepctive
Unifying theme

## Slide 35: Model Mismatch
Mismatch between the data generating mechanism and the model presumed during data fitting is another key source of error in parameter estimates. In qMRI, mismatch is typically due to use of an oversimplified model.
x
f(x)

## Slide 36: Example: Variable Flip Angle T1/R1 Mapping
It is well-known that RF transmit (B1+) is not spatially uniform, especially at high field. If this reality is ignored during model fitting (e.g., universal flip angle), model-data mismatch will cause erroneous T1/R1 values to be generated.
Cheng and Wright. MRM 55:566–574, 2006
Trzasko et al., in preparation
R1 Map – w/o correction
B1+ Map
R1 Map – w/ correction

## Slide 37: Example: Variable Flip Angle T1/R1 Mapping
Keenan et al. MRM 79(1):48–61, 2018
Trzasko et al., in preparation
How significant are errors from ignoring B1+? Even when imaging at 3T with a non-accelerated 5-angle VFA-SPGR sequence and Maximum Likelihood R1 estimation, errors are in the 10-20% range (ISMRM-NIST Phantom):

## Slide 38: Example: Fat+Water Imaging
It is well known that accurate estimation of water+fat from multi-echo GRE data requires fitting with a model that includes off-resonance and R2*. However, ignoring T1, tissue temperature, and concomitant fields can also cause model mismatch and error generation in parameter estimates.
Liu et al., MRM 58(2):354-364, 2007
Hernando et al., MRM 72 (2), 464-470, 2014
Colgan et al., MRM 78 (2), 730-738, 2017

## Slide 39: Example: Simultaneous FWS+MRE
By utilizing a different TE for each motion encoding direction in an MRE scan, chemical shift information is simultaneously encoded. However, the concomitant fields from the motion encoding gradients (MEG) differ by axis. If this effect is not accounted for, water+fat estimates will be erroneous.
Trzasko et al., Proc. ISMRM 2015
Le et al., Proc. ISMRM 2018
Waves - R/L
Waves - A/P
Waves - S/I
Shear Modulus (kPa)
0
20
10
5
15
Water
Fat
B0 Field Map
Reference
Uncorrected
Notes: The stiffness map indicates that vegetable shortening (18.3±4.1kPa) – which is crystalized at room temperature – is substantially more stiff then gelatin

## Slide 40: Example: Simultaneous FWS+MRE
By utilizing a different TE for each motion encoding direction in an MRE scan, chemical shift information is simultaneously encoded. However, the concomitant fields from the motion encoding gradients (MEG) differ by axis. If this effect is not accounted for, water+fat estimates will be erroneous.
Trzasko et al., Proc. ISMRM 2015
Le et al., Proc. ISMRM 2018
Waves - R/L
Waves - A/P
Waves - S/I
Shear Modulus (kPa)
0
20
10
5
15
Water
Fat
B0 Field Map
Reference
Corrected
Notes: The stiffness map indicates that vegetable shortening (18.3±4.1kPa) – which is crystalized at room temperature – is substantially more stiff then gelatin

## Slide 41: Example: Diffusion Imaging
Gradient nonlinearity is typically associated with geometric image distortion. However, it also causes gradient-based motion encoding strength (i.e., b-values) to spatially vary. If this is not accounted for during model fitting, diffusion parameters will exhibit spatially-varying errors.
Tao et al, JMRI, in press
Tan et al., MRM 38(2):448-53, 2013
Bammer et al., MRM 50(3):560-9, 2003
~ isocenter
+6cm (S)
pre-GNC
GNC
pre-GNC
GNC
MD
FA
MD
FA
Difference Ratio Map
Beyond diffusion, gradient nonlinearity affects every motion encoding-based scheme, including flow and elastography.
Notes: The stiffness map indicates that vegetable shortening (18.3±4.1kPa) – which is crystalized at room temperature – is substantially more stiff then gelatin

## Slide 42: Outline
What is Quantitative MRI (qMRI)?
qMRI: a Signal Processing Perspective
Sources of Error in qMRI
  Noise Model
  Signal Model
  Fitting Strategy
Summary and Closing Remarks
Notes: Broaden persepctive
Unifying theme

## Slide 43: Can My Algorithm Really Cause Errors?
x
f(x)
x
f(x)
Convex
Non-Convex
Global Minimum
Local Minimum
In lieu of using general-purpose numerical solvers, qMRI model fitting performance can often be improved by developing/applying custom solvers designed specifically for the target (mathematical) class of problem.
Yes! Many qMRI fitting problems require minimization of non-convex cost functions that contain spurious local minima, whose results may violate data and/or prior assumptions (e.g., phase wrap ≠ smooth).

## Slide 44: Example: Fat+Water Imaging
Hernando et al. MRM 63(1): 79-90, 2010
In fat+water imaging, “swaps” occur if the model fitting algorithm cannot robustly manage cost periodicity or local minima in the ∆B0 estimate. Few algorithms are capable of performing robustly in this scenario – an exception being iterated graph cuts (for binary Markov random fields (MRF)).
Graph Cuts
ICM
Descent
Field Map
Water
Fat

## Slide 45: Example: Fat+Water Imaging
Researchers often rely on vendor-provided qMRI solutions for their studies – always be cognizant of all parts of the tools you adopt, include the fitting algorithms used under-the-hood.
Stinson et al. MRM, 78(6):2203–2215, 2017
mDIXON
MBIR + Graph Cuts

## Slide 46: Example: MR Elastography
We noted earlier that noise can tissue displacement estimation from phase contrast data in MR elastography. For MAP estimators the cost function also exhibits high-dimensional periodicity, yielding a challenging optimization task:
Trzasko et al., Proc. ISMRM 2014
Voxel Solution Space for T=4 Phase offsets
real
imaginary
min
max

## Slide 47: Example: MR Elastography
ML - Descent
ML – Descent - Unwrap
MAP - Descent
MAP – Graph Cuts
Real
Imag
Complex harmonic estimation in rectal MRE; comparison of fitting algorithms:
Trzasko et al., Proc. ISMRM 2014

## Slide 48: Example: MR Elastography
Real
Imag
ML - Descent
ML – Descent - Unwrap
MAP - Descent
MAP – Graph Cuts
Complex harmonic estimation in rectal MRE; comparison of fitting algorithms:
Trzasko et al., Proc. ISMRM 2014

## Slide 49: Back to Error Propagation…
Compared to the propagation of noise-based errors through reconstruction, how algorithmic errors propagate and manifest can be more subtle; e.g., 3D Shells:
Sutton et al. IEEE TMI 22(2):178-188, 2003
Trzasko et al. Proc. RSNA, SSQ17-04, 2012
Tao et al. Medical Physics 42(12):7190–7201, 2015.
Field Map
Gridding Recon
None
MAP - Descent
MAP - Robust

## Slide 50: Back to Error Propagation…
Compared to the propagation of noise-based errors through reconstruction, how algorithmic errors propagate and manifest can be more subtle; e.g., 3D Shells:
Sutton et al. IEEE TMI 22(2):178-188, 2003
Trzasko et al. Proc. RSNA, SSQ17-04, 2012
Tao et al. Medical Physics 42(12):7190–7201, 2015.
Field Map
Gridding Recon
None
MAP - Descent
MAP - Robust

## Slide 51: Summary
  In this lecture we:
  Defined and overviewed qMRI
  Discussed how qMRI processing is fundamentally a parametric model fitting problem
  Introduced statistical signal processing tools that are useful for qMRI applications
  Defined and demonstrated several independent sources error in qMRI processing
  Described modern signal processing tools for mitigating qMRI error
Notes: Take on old JAVA phrase

## Slide 52: Closing Remarks
  Hopefully you now have an appreciation for the many potential sources of error in qMRI parameter estimation, and the tools that can be used to both characterize and mitigate them.
  Whether you’re revisiting an old qMRI methods or checking out a new one, approach all with the same rigor and scrutiny.
Notes: Take on old JAVA phrase

## Slide 53
?
trzasko.joshua@mayo.edu